## Slide 1: Introduction to Pathology Informatics
Session 5: Digital Imaging in Pathology
James Harrison, MD PhD
University of Virginia
2023
API distribution, version 1.0
Notes: These 10 slide sets and term lists are being provided through the Association for Pathology Informatics (API) with a license that allows free use, modification, and redistribution with attribution (see the permitted uses slide at the end of each presentation). The design has been kept relatively basic to make editing and adaptation easier. Users are encouraged to modify and improve the slides, share the changes (with attribution of the original material), and possibly contribute the improved slide sets back to the API or other distributors. As designed, the slides are relatively wordy so that they are easier to follow without an instructor’s guide (though one may appear in the future). Feel free to edit the slides to streamline them and optimize them for local teaching style and available time. Some sessions contain optional slides with extra detail, and these are indicated in the slide notes. Some sessions contain slide builds (most are simple) that don’t print well; these sessions have optional slides for printing, and these are indicated in the slide notes.
The goal of these slide sets is to provide foundational information that will allow pathology trainees to more effectively learn and practice pathology, not to teach the performance of specific tasks in pathology. The overall learning objectives are listed on the third slide of session 0 with some explanatory comments in the notes. Subsequent sessions have agendas but not separate learning objectives because the learning objectives remain the same.
The slides have been used with a teaching schedule in which the orientation is given during the summer and then the remaining sessions are give once a month (with lunch) from September to May. They could also fit into a 2, 4, or 5-6 week teaching module if taught on a weekly/semiweekly basis.

## Slide 2: Why the interest in digital pathology now?
Radiology went digital ~20 years ago. Why?
What are the compelling incentives for digital pathology?
  Not lower cost, reduced process complexity, faster processing (yet)
  Remote diagnosis when availability of pathologists is limited
  New applications and tools for diagnosis
    Decision support for pathologists with improved accuracy and consistency
    Efficiency and new metrics: counting objects, identifying objects/regions of interest
    Machine learning
  Improved workflow and access to slide images
  Archival images when slides need to be repurposed
  Clinical consultation (pathology, tumor board)
  Education

## Slide 3: Digital imaging agenda
Types and characteristics of digital images
Image compression
DICOM data standards for medical imaging
Digital pathology system details
Instrument/system components for digital pathology
  Features
  Interconnection (interfaces)
Calibrating and standardizing digital pathology
Validation and approval of digital pathology systems
Agenda icon created by Freepik – Flaticon

## Slide 4: Types of pathology digital imaging systems
Telepathology (relatively well-established)
  Microscope with integrated camera for static or live streaming images
  Robotic stage/focus/lenses allows remote operation by a pathologist
  Data may or may not be captured for storage, bandwidth requirements low
  May include low-capacity multi-slide loaders (2-5 slides)
  Remote pathology, frozen section, cytology
Whole slide imaging (limited installation for clinical diagnosis so far)
  Digitized image of all tissue on a slide
  May be used for local or remote image viewing
    Expensive, requires specialized equipment
    Bandwidth requirements may be high

## Slide 5: Types of digital images
Mathematical shapes
  Max resolution at any size
  Smaller files
Solid colors or defined color gradients
  Graphic art
“Painted” or “bitmapped” images
Array of dots (pixels)
  Visible with enlargement
  Large files if high resolution
  Anti-aliasing
Complex images
  Photographs
  Pathology images
Vector
Raster
Image from Wikimedia Commons
File types:
SVG, EPS, CGM, AI, CDR, PDF
BMP, RAW, TIFF, GIF, PNG, JPEG
(PDF can contain both)

## Slide 6: The colored Pixel (Picture Element)
Red
0-255
Green
0-255
Blue
0-255
102
(40%)
26
(10%)
128
(50%)
RGB Colors
Color data (3 bytes)
Other data optional
Transparency, etc.
RGB: (102, 26, 128)
One byte for each color

## Slide 7: Characteristics of raster images
Resolution
  More pixels = more accurate representation
  Image resolution: size in pixels (1280 x 1024)
  Optical resolution: Distance between structures that can be seen distinctly
    In microscopy, related to numerical aperture of lens and wavelength
  Digital microscopy: Both are important and should be complementary
    Size of pixels is defined by the camera sensor size and magnification
    20x equivalent ~ 0.5 µm per pixel; 40x equivalent ~ 0.25 µm per pixel.
Color depth (“bit depth”), a form of resolution
  Full color is 24-bit (8 bits or 256 levels of red, green, & blue)
Aspect ratio
  Shape of image (wide, square, tall)
High px density
Med px density
Low px density
Image from Wikimedia Commons

## Slide 8: Pathology images: View and movement
When the image is larger than the view
  From photography/cinematography
Field of view (FOV), what’s visible
FOV movement
  Pan in X and Y directions
  Zoom (magnification)
  “Z-stacking,” vertical with respect to slide
    Multiple images at different vertical focus levels
    Up to 30 layers for focus level (cytology) and 3D reconstruction

## Slide 9: Image data size
Directly related to resolution: (color depth) x (# pixels)
  (color depth) x (tissue size) / (pixel size @ magnification)
  24 x (1 mm)2 / (0.25 µm)2 = 384,000,000 bits = 48 MB per square mm
Z-stacking? Multiply image size x number of levels
File compression is necessary for digital pathology
Even with compression, smart use of storage and network resources is required

## Slide 10: Image compression
Lossless (TIFF, PNG), best for line/graphic art
Lossy (TIFF, JPEG) best for photographs & digital pathology
  Variable compression with artifacts; do not expand/edit/recompress
  Represents small blocks of pixels as functions, reconstruction isn’t perfect
  JPEG, discrete cosine transform; JPEG2000, wavelet transform (preferred)
Size reduction
A, C: 16-fold
B, D: 100-fold
Zarella MD, Bowman; D, Aeffner F, et al. A practical guide to whole slide imaging. Arch Pathol Lab Med. 2019;143(2):222-234. doi:10.5858/arpa.2018-0343-RA
Notes: Compress and safe as final step. Don’t expand, edit, recompress.

## Slide 11: Image files (simple example)
TIFF: Tagged Image File Format (baseline)
8-byte header (identifies as TIFF, gives location of first image file directory)
  TIFF files may contain one or several images (subfiles)
Image file directory (list of “tags” with values for each image)
  Each entry is 12 bytes (2-byte tag number and 10 bytes value): encodes image characteristics
  Many tags available, for example:
    Type of image (B&W, grayscale, indexed color, full color)
    Number of samples per pixel (“full color” is 3 samples/px, 8-bits each for R/G/B)
    Resolution & aspect ratio
    Compression scheme, if any
    Starting locations of rectangular blocks of data (strips) that make up the image
Image data (divided into strips)
IFD
Strip 1
Strip 2
Strip 3
Strip 4
Header
Simplified TIFF file structure

## Slide 12: “Image pyramid” in digital pathology
Scan at high resolution
Pre-compute and store subsampled images at lower resolution
Send smaller, lower resolution to client when possible
Send only FOV of higher resolution images
Supported by image database
Multiple levels
Wang Y, Williamson KE, Kelly PJ, James JA, Hamilton PW (2012) SurfaceSlide: A multitouch digital pathology platform. PLOS ONE 7(1): e30783. https://doi.org/10.1371/journal.pone.0030783

## Slide 13: Data standard for medical imaging (DICOM)
Digital Imaging and Communications for Medicine
File formats and messaging for medical imaging
Extensive application in radiology
Developing for pathology
Herrmann M, Clunie D, Fedorov A, et al. Implementing the DICOM standard for digital pathology. J Pathol Inform. 2018;9(1):37. doi:10.4103/jpi.jpi_42_18

## Slide 14: Encoding the image pyramid
Each image in the pyramid is a separate DICOM data set
Each image is divided into tiles which are encoded as compressed frames
The header information for each image contains the size and location of each frame
Frames are individually accessible in the image database
Herrmann, et al. J Pathol Inform. 2018;9(1):37. https://doi.org/10.4103/jpi.jpi_42_18

## Slide 15: Accessing the pyramid
Client software for image review and signout; DICOM communications
Image access by frame to support the current magnification & FOV
Tools for image organization & measurement, annotation, syncing multiple images, object ID & counting, segmentation, etc.
MSK Viewer images provided by Matthew Hanna MD

## Slide 16: Accessing the pyramid
Client software for image review and signout; DICOM communications
Image access by frame to support the current magnification & FOV
Tools for image organization & measurement, syncing multiple images, object ID & counting, segmentation, etc.
MSK Viewer images provided by Matthew Hanna MD
Notes: Static slide for printing or PDF. Hide for presentation.

## Slide 17: Digital pathology workflow
Modified from Digital Pathology Association. Interoperability between Anatomic Pathology Laboratory Information Systems and Digital Pathology Systems.
https://digitalpathologyassociation.org/_data/cms_files/files/Interoperability_Between_Anatomic_Pathology_Laboratory_Information_Systems_and_Digital_Pathology_Systems.pdf
(*Remote or local)
PACS
Communications relatively proprietary so far
Notes: PACS: Picture Archiving and Communications System

## Slide 18: Standards for inter-system communication • IHE Pathology and Laboratory Medicine (PaLM) Workgroup intersystem messaging • HL7 messaging standard for digital pathology in development
Order entry
LIS
Result display
Slide scanner
Human or algorithmic
interpretation
Dash RC, Jones N, Merrick R, et al. Integrating the Health-care Enterprise pathology and laboratory medicine guideline for digital pathology interoperability. J Pathol Inform. 2021;12:16. doi: 10.4103/jpi.jpi_98_20
Actors and Roles:
PACS
IHE Profile: A standard set of messages for (eg):
Worklist exchange
Scanned image exchange
Image display
Image interpretation

## Slide 19: Scanner details
Slide feed, light source, stage, lens, high resolution camera, data processing & communications
Appropriate color gamut
  Compatible with display
Minimize input artifacts
  Dirt, dust, bubbles, fingerprints, markings
Scanner performance must be validated
  Before putting into service
  Periodically to identify performance change (performance monitoring)
Patel A, Balis UGJ, Cheng J, et al. Contemporary whole slide imaging devices and their applications within the modern pathology department: A selected hardware review. J Pathol Inform. 2021;12:50. doi: 10.4103/jpi.jpi_66_21
Image from Wikimedia Commons
Color gamut curves

## Slide 20: Slide imaging
Area (tiles) or line capture
Tissue detection & focus pattern
Brightfield/fluorescence/ multispectral
20x, 40x, oil on some
Auto-feed, batch sizes, slide sizes
Image quality
  Color accuracy/balance, brightness, contrast
  Compression, focus, stitching, clipping
Leica Aperio GT 450
450 slides on stainer racks
81 slides per hr at 40x (1.5x1.5 cm)
0.26 µm/pixel at 40x
1 Gb/s to image server
Leica Biosystems, used with permission
Patel A, Balis UGJ, Cheng J, et al. J Pathol Inform. 2021;12:50.
Notes: TDI = time, delay, integration; can be used for fluorescence; Need to get permission to use the Aperio scanner image, or replace.

## Slide 21: Example imaging problems
Overcompression1
Stitching2
Contamination2
Focus1
Pantanowitz L, Liu C, Huang Y, Guo H, Ronde GK. Impact of altering various image parameters on human epidermal growth factor receptor 2 image analysis data quality. J Pathol Inform. 2017;8(1):39. doi:10.4103/jpi.jpi_46_17
Patel AU, Shaker N, Erck S, et al. Types and frequency of whole slide imaging scan failures in a clinical high throughput digital pathology scanning laboratory. J Pathol Inform. 2022;13:100112. doi:10.1016/j.jpi.2022.100112

## Slide 22: PACS details
May be local or SaaS (cloud)
Communicates with scanner for image data (~1 Gb/s)
Communicates with LIS for patient data and reporting
Communicates with display for presentation of images and signout (at least 20 Mb/s, up to 10 Gb/s)
Provides an image database
  Special purpose; hierarchical, network, or object-oriented
  High volume data, typically RAID hardware for speed/reliability
  Indexed by textual metadata, organized by case/series/image
Supports access to partial image data (saves bandwidth)
  Implements DICOM standards including the image pyramid
  Understands header and pixel data, strips/tiles/frames of pixels
Database icon created by phatplus – Flaticon
Histology from MSK, radiology from Wikimedia Commons
DICOM

## Slide 23: Display details
High resolution, adequate color gamut
Minimum PACS connection bandwidth recommended 20 Mbit/s or more
Display may be proprietary to scanner and/or PACS, or required by FDA approval
Tools for:
  Worklists
  Case/slide/image retrieval and organization
  Navigation within images; pan/zoom
  Alignment/overlay of images
  Measurement and counting of objects
  Image and region identification (AI)
  No standard toolkit yet, may be a la carte
MSK Viewer images provided by Matthew Hanna MD

## Slide 24: Analytics software likely functions
Identify tissue
Identify, annotate, and measure regions of interest
Manually or automatically count objects in field or ROI
Classify and annotate cell types
Images from QuPath documentation (CC-BY)

## Slide 25: Calibration and standardization
Optimize imaging via process standardization and best practices
  Histology procedure and reagents
  Type, configuration, and settings of scanners
  Data compression algorithm and degree of compression
  Viewing application, type and age of displays
Color optimization
  Suboptimal color/contrast makes pathologist work more difficult and slower
  Calibration so that scanner color gamut matches display color gamut
    International Color Consortium (ICC) methods using color calibration slides, but may reduce total color gamut, contrast, and luminosity
  Alternatively, optimize contrast and color visibility rather than accuracy
    Particularly useful to improve contrast in fluorescence microscopy with low light
Notes: ICC calibration can reduce total color gamut, contrast, and luminosity.

## Slide 26: Validation of digital pathology systems (CAP)
Strong recommendations
Validation should include 60 cases per application (eg, primary dx with H&E), and 20 cases each for related applications such as IHC.
There should be a 2-week washout period between glass and digital slide evaluations
Concordance should be judged for the same observer, less than 95% should be investigated
Good practices
Labs should validate WSI systems
Validation should be relevant to the expected case types and usage context
Validation should emulate real workflows
Validation slides should be viewed in random order
The entire WSI system should be validated
Procedures should be in place to deal with WSI system changes, including re-validation.
Pathologists should be adequately trained on the system and involved in the validation
Digital images should be confirmed to contain all diagnostic material on slides
Validation documentation should be maintained
Evans AJ, Brown RW, Bui MM, et al. Validating whole slide imaging systems for diagnostic purposes in pathology: Guideline update From the College of American Pathologists in collaboration with the American Society for Clinical Pathology and the Association for Pathology Informatics. Arch Pathol Lab Med. Published online May 18, 2021. doi:10.5858/arpa.2020-0723-CP

## Slide 27: Regulatory review of digital pathology systems
Systems can be validated by the lab as lab-developed tests
Specific software, scanner, and display device FDA-approved as a unit
  Philips Intellipath, 2017
  Leica Aperio, 2019
  Sectra, 2020
First FDA approval of software for use with multiple scanners
  Inspirata Dynamyx, March 2022

## Slide 28: Key points
Digital pathology is on the cusp, pushed by efficiency, consistency, and analytics goals
  Early implementations are in progress, likely to spread if tools are compelling
There are a variety of new technologies and concepts unfamiliar to most pathologists (knowledge growth opportunity  )
Quality management and performance verification will be important
The details of digital tools and their regulation are still developing
  The business case for DP investments is tied to the utility of the tools
  Will be interesting for pathologists to follow and contribute
Data standards are incomplete and in active development
Rate of FDA approvals will increase over the next several years
Key icon created by juicy_fish – Flaticon
Notes: QA and PV is familiar to most pathologists

## Slide 29: Readings
Harrison J. Management of Pathology Information Systems, Chapter 6. In: Laboratory Administration for Pathologists. 2nd ed. CAP Press; 2019:286. pp 95-96, Data Representation, Image Compression. Available for download from the Pathology Informatics Essentials for Residents curriculum page.
Hanna MG, Ardon O, Reuter VE, et al. Integrating digital pathology into clinical practice. Mod Pathol. 2022;35(2):152-164. doi:10.1038/s41379-021-00929-0
Additional resources for reference
Patel AU, Shaker N, Erck S, et al. Types and frequency of whole slide imaging scan failures in a clinical high throughput digital pathology scanning laboratory. J Pathol Inform. 2022;13:100112. doi:10.1016/j.jpi.2022.100112
Evans AJ, Brown RW, Bui MM, et al. Validating whole slide imaging systems for diagnostic purposes in pathology: Guideline update from the College of American Pathologists in collaboration with the American Society for Clinical Pathology and the Association for Pathology Informatics. Arch Pathol Lab Med. 2021;146(4):440-450. doi:10.5858/arpa.2020-0723-CP
Patel A, Balis UGJ, Cheng J, et al. Contemporary whole slide imaging devices and their applications within the modern pathology department: A selected hardware review. J Pathol Inform. 2021;12(1):50. doi:10.4103/jpi.jpi_66_21
Jahn SW, Plass M, Moinfar F. Digital pathology: Advantages, limitations and emerging perspectives. Journal of Clinical Medicine. 2020;9(11):3697. doi:10.3390/jcm9113697
Book icon created by Good Ware – Flaticon

## Slide 30: Permitted uses
© 2023 James H. Harrison, Jr. except as noted below. The original material in this slide set is made available under a Creative Commons Attribution 4.0 (CC BY 4.0) International License. The contents may be used without restriction, modified, and redistributed if they are attributed and any changes are indicated. Some graphical elements are used with permission or according to their own licenses requiring attribution. Attribution links are provided on the slide of use or on the Permitted Uses slide (this slide). These elements may be used and included in modified and redistributed work consistent with the licenses specified at the linked sites, provided that the links are included in the redistributed material.
The cover image is a composite of two laboratory images from Pixabay available for use without attribution.
Some figures are reproduced from The Journal of Pathology Informatics and PLOS One which are open access journals. Several figures are from articles by Zarella et al. and Evans et al. in the Archives of Pathology and Laboratory Medicine, and those articles are open access. See the journal sites at the attribution links provided for the applicable licenses.